## Slide 1: FITPix – Timepix Test D07-0108
Mario Iván MARTÍNEZ
Eduardo MORENO

## Slide 2: First chip settings
Bias Voltage (we used 100 V)
Digital test of pixels
DAC’s scan

## Slide 3: Threshold (THL) equalization
Timepix amp. discrimination level can be chosen to filter out noise.
A single THL level can be set common to all 65,536 pixels
  One wants to put this above the noise level. Typically 40 THL counts ≈ 100 e-
Every pixel has a different noise level
  Gaussian distributed
  4-bit DAC to adjust the “noise” (actually, the amp’s output DC level) level individually
Pixelman feature exists to automatically find the noise level and adjust this DAC for each pixel

## Slide 4: Noise distribution before and after THL equalization
Highest DAC adjust. mean = 444.46 THLcounts std. dev. = 9.67
Lowest DAC adjust. mean = 368.58 THLcounts std. dev. = 9.58
Equalized adjust. mean = 405.53 THLcounts std. dev. = 1.25

## Slide 5: DAC’s adjustment distribution after equalization

## Slide 6: Materials
Timepix chip D0-W0108
Radioactive sources
  Iron: 55Fe
  Americium: 241Am
  Strontium: 90Sr
Software
  MAFalda v 2.1
  Pixelman 2.1.1 (7.12.2011)
Hardware
  FITPix (Medipix/TimePix USB 2.0 v.2.3, Firmware 3)
  Timepix (D07-W0108)
    300 m thick

## Slide 7: Acquisition parameters
20o C
10 milliseconds per frame
100,000 frames
Operation mode Time Over Threshold
Bias voltage 100V
Distance source-sensor 2 centimeters

## Slide 8: Tested on
Desktop, Windows 7
Laptop, Windows 7
Laptop, Mac
Virtual Machine, Linux
Problems with usb connection in Mac and Virtual machine

## Slide 9: Measurements

## Slide 10: Contrast
To observe pixels’ noise (individuate noisy ones)
100,000 frames
Acquisition mode Time Over Threshold
Time per frame 10 milliseconds

## Slide 11
High noise (bad bonding?)

## Slide 12: Contrast

## Slide 13
zoom for each “quadrant”

## Slide 14: zoom to inner “good” area

## Slide 15: Clusters Single Double Triple Quad

## Slide 16: 90Fe

## Slide 17: ToT distribution of single-pixel clusters

## Slide 18: ToT distribution of single- and double-pixel clusters

## Slide 19: Number of clusters vs. total ToT per frame

## Slide 20: 241 Am

## Slide 21: Single-pixel clusters

## Slide 22: ToT distributions for different cluster sizes

## Slide 23: Number of clusters vs. total ToT per frame

## Slide 24: Conversion Energy to ToT

## Slide 25: Aplicaciones a física médica
Imágenes bidimensionales con rayos X
Imágenes tomográficas con rayos X
Dosimetría con radiación dispersa
Estudio de fragmentación de núcleos en hadroterapia

## Slide 26: Medipix 3 TSV layout
Complete map
Metrology boxes
Front side Electrical tests area
Back side Electrical tests area

## Slide 27
Modular unit for imaging applications:
NxM matrix of detector tiles in a
module
NxM matrix of modules
BGA bonded to a ceramic
multilayre board